- ASTRA 6 Report PSty standard_2-4-15

File Name: C:\Documents and Settings\Ayres Research Group\My Documents\Emily\Other's research\PSty standard_2-4-15.afe6 
Collection Operator: AYRESGPC\Ayres Research Group (AYRESGPC\Ayres Research Group (Ayres Research Group)) 
Processing Operator: AYRESGPC\Ayres Research Group (Ayres Research Group)

- Sample: sample

Concentration: 0.000 mg/mL

- Configuration

Concentration Source: RI 
Flow Rate: 1.000 mL/min

- Light Scattering Instrument: miniDAWN TREOS

Band Broadening Correction: Yes (Instrumental: 37.238 μL, Mixing: 76.699 μL) 
Cell Type: Fused Silica 
Wavelength: 656.0 nm 
Calibration Constant: 6.1272×10-5 1/(V cm)

| Detector | Scattering angle | Gain | Normalization coefficient |
| --- | --- | --- | --- |
| 1 | 46.6º | n/a | 0.764 |
| 2 | 90.0º | n/a | 1.000 |
| 3 | 133.4º | n/a | 0.810 |

- RI Instrument: Optilab rEX

Band Broadening Correction: Yes (Instrumental: 27.923 μL, Mixing: 0.732 μL) 
Wavelength: 658.0 nm

- Viscometer: ViscoStar

Band Broadening Correction: n/a 
Dilution Factor: 0.4976

- Solvent: thf

Refractive Index: 1.402

| Fluid Connections |  |  |
| --- | --- | --- |
| Source Instrument | Destination Instrument | Delay Volume (mL) |
| Generic Pump | Injector | 0.000 |
| Injector | Generic Column | 0.000 |
| Generic Column | miniDAWN TREOS | 0.000 |
| miniDAWN TREOS | ViscoStar | 0.359 |
| ViscoStar | Optilab rEX | 0.107 |

- Processing

Collection Time: Wednesday February 04, 2015 03:38:45 PM Eastern Daylight Time 
Processing Time: Wednesday February 04, 2015 04:13:56 PM Eastern Daylight Time

- Basic Collection:

LS Instrument Collection Interval: 1.000 sec

- Baselines:

| Series | Start | Stop | Type |
| --- | --- | --- | --- |
| detector 1 | (14.297, 0.044) | (27.927, 0.043) | manual x, auto y |
| detector 2 | (13.243, 0.029) | (26.486, 0.029) | manual x, auto y |
| detector 3 | (11.136, 0.043) | (27.575, 0.043) | manual x, auto y |
| differential refractive index data | (13.501, -0.000) | (22.904, -0.000) | manual x, auto y |
| Differential Pressure | (9.072, -0.001) | (27.913, -0.001) | manual x, auto y |

- Peak settings:

| Peak Name | Peak 1 |
| --- | --- |
| Light Scattering Model | Zimm |
| Fit Degree | 1 |
| dn/dc (mL/g) | 0.1821 |
| A2 (mol mL/g²) | 0.000 |
| Viscometry Model | Huggins |
| Huggins Equation Parameter | 0 |
| Kraemers Equation Parameter | 0 |

- Results Fitting Procedure:

| Data | Fit Model | Degree | R2 | Extrapolation |
| --- | --- | --- | --- | --- |

- Results

| Peak Results |  |
| --- | --- |
|  | Peak 1 |
| Hydrodynamic radius (v) moments (nm) |  |
| Rh(v)n | 9.288 (±0.170%) |
| Rh(v)w | 9.521 (±0.164%) |
| Rh(v)z | 9.735 (±0.166%) |
| Rh(v)(avg) | 9.963 (±0.014%) |
| Masses |  |
| Calculated Mass (µg) | 649.84 |
| Molar mass moments (g/mol) |  |
| Mn | 1.109×105 (±0.438%) |
| Mp | 1.118×105 (±0.386%) |
| Mv | 1.150×105 (±0.044%) |
| Mw | 1.148×105 (±0.401%) |
| Mz | 1.185×105 (±0.895%) |
| Mz+1 | 1.223×105 (±1.440%) |
| M(avg) | 1.363×105 (±0.021%) |
| Polydispersity |  |
| Mw/Mn | 1.035 (±0.594%) |
| Mz/Mn | 1.069 (±0.996%) |
| rms radius moments (nm) |  |
| Rn | 15.2 (±6.7%) |
| Rw | 15.5 (±6.1%) |
| Rz | 15.7 (±5.8%) |
| R(avg) | 19.5 (±0.2%) |
| Intrinsic viscosity moments (mL/g) |  |
| [η]n | 46.224 (±0.221%) |
| [η]w | 48.00 (±0.22%) |
| [η]z | 49.652 (±0.243%) |
| [η](avg) | 47.240 (±0.014%) |

Mark-Houwink-Sakurada a: 1.117 (±0.095%) 
Mark-Houwink-Sakurada K: 1.074×10-4 (±1.243%) mL/g