CHROMOPHORIC POLYMERS

CHARACTERIZED BY MALS DETECTION

Michelle H. Chen

Wyatt Technology Corporation

Abstract

The use of forward laser monitor in a multi-angle light scattering (MALS) detector, such as DAWN or miniDAWN, has been established for the correction of light scattering measurements of polymers absorbing light at the operating wavelength. In this poster such corrections are demonstrated in both micro-batch and chromatography modes using a flow cell.

P o l y m e r

Chromophoric?

Yes*						  No	        MALS

Fluorescence			Absorption

Add filters to

light scattering

photo-diodes

Correct by using forward laser monitor

* Generally speaking, fluorescence and absorption will only affect the determination of molar mass not the root-mean square radius in light scattering measurements.

Theory

L

Laser, I0

Forward Laser monitor, F

when no absorption, F0

Intensity of the laser at the center

I true

Beer-Lambert Law:

-log(F/F0) = eLc

I true = I0 F/F0

-log(Itrue/ I0) = eLc/2

where e is the extinction coefficient of the polymer at the laser wavelength and c is the concentration.

Micro-batch Analysis of a Colored Protein

C1

C2

C3

Solvent

Solvent

Experimental conditions: solvent, PBS with surfactant; detector, DAWN DSP; sample, Bacterial Reaction Center (a protein complex); concentration, C1 < C2 < C3.

Output of the forward laser monitor signal was collected via an auxiliary channel.

To obtain corrected intensity from the sample:

(ILS (sample) - ILS (solvent) F/F0) /F/F0

or

ILS (sample) /F/F0- ILS (solvent)

Note that the correction factor of F/F0 is for flow cell, the factor is F/F0 for scintillation vial.

Using ASTRA  to Make Corrections for Absorption

Step 1

Rescale the baselines by multiplying each voltage with F/F0.

Step 2

Calculate the apparent MM in Debye plot and divide it by F/F0 to obtain the true MM.

or

Export the data and calculate the results in a spreadsheet.

Micro-batch Analysis of a Colored Protein

C1

C2

C3

Solvent

Solvent

Experimental conditions: solvent, PBS with surfactant; detector, DAWN DSP; sample, Bacterial Reaction Center (a protein complex); concentration, C1 < C2 < C3.

Output of the forward laser monitor signal was collected via an auxiliary channel.

Size Exclusion Chromatography of a Colored Protein

LS at 90°

RI

F monitor

Volume (mL)

Experimental conditions: column, TSK G3000SWxL; mobile phase, PBS; detectors, DAWN DSP and Optilab; sample, bovine liver catalase.

Molar Masses Determined with and without Corrections for Absorption

RI trace of BLC tetramer peak

uncorrected

corrected

Molar masses for Bovine Liver Catalase (BLC) tetramer:

Theoretical, 240 kD; uncorrected, 220 kD; corrected, 239 kD.

Conclusions

- Forward laser monitor can be used to correct for absorption by absorbing polymers at the laser wavelength.

- The correction is only needed for molar mass determination, and generally not necessary for size measurement.

- The accuracy of the measurement will deteriorate when the attenuation of the laser intensity is severe.

- The use of a higher wavelength laser may eliminate absorption and fluorescence.

Appendix: Use ASTRA and Excel to perform the absorption correction

- Use ASTRA to calculate the apparent MW* of a small non-absorbing polymer (like the one for normalization) with the signal from only one detector at a particular concentration, C*. Record the MW*, C*, and voltage V*(baseline subtracted).

- Export the raw voltages to Excel to obtain the following voltage at each LS detector you want to use

ILS (sample) /F/F0- ILS (solvent)

- Plot voltages at each concentration, C, as a function of sin2(thetha/2) to obtain the voltage at zero scattering angle, V.

- The correct MW = MW* x C* x V/(C x V*).